CAMAC TRIGGER GENERATOR for INO, INO_CTG

Contents: 

· Specifications
· User manual
· Test setup and observations
· Application note
· Technical reference manual

CAMAC TRIGGER GENERATOR for INO, INO_CTG

Specifications 

The CAMAC Trigger Generator, INO_CTG has been designed to generate final trigger based on Predefined Coincidence Criterion on a large number of input signals from the prototype Neutrino detector. Additionally it has SCALARS for the FINAL_TRIGGER and for a number of intermediate coincidence logic signals. The SCALARS can be read on CAMAC. The valid signals are latched with final trigger and can also be read on CAMAC. Inputs can be individually deselected from taking part in the trigger generation. 

It has been implemented as a two width CAMAC module and most of the logic has been implemented in Altera 10K50 field programmable gate array device.


This module accepts first level trigger information derived from the individual layers of the detectors, namely in terms of one, two, three and four fold signals. Further COINCIDENCE LOGIC is implemented in the INO_CTG module. There are a total of 112 signals coming on LVDS lines to the CAMAC module. 

Trigger signals generated from the detector assembly:

Input - 
One, two, three and four fold signals each for 14 X and 14 Y layers




Total 28*4 = 112 from Level 1


Signals -  
Low Voltage Differential Signaling (LVDS) with 400mV





differential signals 

Connectors - 
4x60 pin FRC standard connectors with 0.1” pitch




28 signals on each connector

Inputs to be connected on twisted pair flat cable.

Output - 
LTO - Logic trigger out, fanout of 1 as NIM output

FTO - Final trigger out = LTO & not LAM

fanout of 3 as NIM output

LTO & FTO width in multiples of 50 ns (1 to 7 clocks, default 50 ns)

LAM – CAMAC Look At Me generated with FTO

Multiplicity Register – All 112 inputs latched on FTO, to be read 

on CAMAC 




Scalers – Eight scalers on intermediate triggers, one on LAM width
CAMAC TRIGGER GENERATOR for INO, INO_CTG

User_manual

This module has been implemented as a two width CAMAC module and most of the logic has been implemented in Altera 10K50 field programmable gate array device. Low Voltage Differential Signaling (LVDS) has been used as input interface signal standard. Output signals are generated as NIM logic level. It has been designed to generate final trigger based on Predefined Coincidence Criterion on a large number of input signals from the prototype Neutrino detector. Additionally it has SCALARS for the FINAL_TRIGGER and for a number of intermediate coincidence logic signals. The SCALARS can be read on CAMAC. The valid signals are latched with final trigger and can also be read on CAMAC. Inputs can be individually deselected from taking part in the trigger generation. 

The individual blocks of the module are described in details.

Trigger Logic: 

The primary functionality of the module is to generate a trigger based on pre-defined trigger pattern on 112 inputs. It can be described with the help of a block diagram as follows:

 
[image: image1.png]Two Fold Inputs. 14X

Two Fold Inputs. 14Y

Two Fold Coincedence Logic Block

j>_

.,
D
o
D

—P SCALAR_TO_F2%X

| SCALAR_TO_F2Y




 
[image: image2.png]Three Fold Inputs. 14X

Thiee Fold Inputs. 14Y

Three Fold Coincedence Logic Block

j>_

B
o)
B
D

—P SCALAR_TO_F3X

| SCALAR_TO_F3Y





[image: image3.png]Four Fold Coincedence Logic Block

Four Fold Inputs. 14X

—P SCALAR_TO_F4X

Four Fold Inputs.14Y —p SCALAR_TO_F4Y




 

Trigger Logic in Boolean equations: 


TO_F1X 
= 
F1LX1 & F1LX2 & F1LX3 & F1LX4  & F1LX5

      + 
F1LX2 & F1LX3 & F1LX4 & F1LX5  & F1LX6

      + 
F1LX3 & F1LX4 & F1LX5 & F1LX6 & F1LX7


:


:

      + 
F1LX10 & F1LX11 & F1LX12 & F1LX13 & F1LX14



TO_F1Y 
= 
F1LY1 & F1LY2 & F1LY3 & F1LY4 & F1LY5 

      + 
F1LY2 & F1LY3 & F1LY4 & F1LY5  & F1LY6

      + 
F1LY3 & F1LY4 & F1LY5 & F1LY6 & F1LY7


:


:

      + 
F1LY5 & F1LY11 & F1LY12 & F1LY13 & F1LY14



TO_F2X 
= 
F2LX1 & F2LX2 & F2LX3  & F2LX4

      + 
F2LX2 & F2LX3 & F2LX4  & F2LX5

      + 
F2LX3 & F2LX4 & F2LX5  & F2LX6


:


:

      + 
F2LX11 & LX12 & F2LX13 & F2LX14



TO_F2Y 
= 
F2LY1 & F2LY2 & F2LY3  & F2LY4

      + 
F2LY2 & F2LY3 & F2LY4 & F2LY5

      + 
F2LY3 & F2LY4 & F2LY5  & F2LY6


:


:

      + 
F2LY11 & F2LY12 & F2LY13 & F2LY14



TO_F3X
= 
F3LX1 & F3LX2  & F3LX3

      + 
F3LX2 & F3LX3  & F3LX4

      + 
F3LX3 & F3LX4 & F3LX5


:


:

      + 
 F3LX12 & X13 & F3LX14



TO_F3Y
=
F3LY1 & F3LY2  & F3LY3

      + 
F3LY2 & F3LY3  & F3LY4

      + 
F3LY3 & F3LY4 & F3LY5


:


:

      + 
F3LY12 & F1LY13 & F3LY14



TO_F4X
= 
F4LX1   & F4LX2

      + 
F4LX2   & F4LX3

      + 
F4LX3  & F4LX4


:


:

      + 
 F4LX13 & F4LX14



TO_F4Y
= 
F4LY1   & F4LY2

      + 
F4LY2   & F4LY3

      + 
F4LY3  & F4LY4


:


:

      + 
 F4LY13 & F4LY14

LTO

= 
TO_F1X + TO_F1Y 

    +
TO_F2X + TO_F2Y

    +
TO_F3X + TO_F3Y

    + 
TO_F4X + TO_F4Y

Note: LTO is assigned to one of the front panel NIM outputs on Lemo connector

LAM_ENABLE.SET = N & A (0) & F (26)

LAM_ENABLE.RESET  = N & A (0) & F (24) +Z S2

LAM_REQ.D = 1

LAM_REQ.CLK = /FTO

LAM_REQ.RESET = N & A (0) & F (10) + ZS2

LAM = LAM_REQ & LAM_ENABLE

FTO  = LTO & /LAM

Note: FTO is assigned to three of the front panel NIM outputs on Lemo connectors

Q = LAM  (CAMAC command N & A (0) & F (8))

LAM LOGIC: 

Final LAM signal is sent on CAMAC bus when FTO is generated and LAM is enabled through CAMAC command N & A(0) & F(26). LAM can be disabled through CAMAC command N & A(0) & F(24) and reset through CAMAC command N & A 0) & F(10). Once LAM is generated and till it is cleared,  another event will generate only LTO and not FTO.  

[image: image4.png]LAM_mask
LAM disable —{Preset @
F(24)A(0)
LAM enable —|Clear
F(26)A(0) L
1
Lo
q
LAM_re
Fo —| e
Clear
Reset LAM

F(10)A(0) or Z

CAMAC
LAM




MASK logic
Mask registers are provided to individually enable / disable inputs from taking part in the coincidence. After power-on all channels are masked. According to the bits in the mask corresponding inputs will be masked (bit =0) or unmasked (bit = 1). The Mask registers are 8 bit wide and the bits assignment is as follows:

F1X8 – F1X1 mask register 

N & A(0) & F(17) with data bits as shown: 

	Bit 7

F1LX8
	Bit  6

F1LX7
	Bit 5

F1LX6
	Bit  4

F1LX5
	Bit  3

F1LX4
	Bit 2

F1LX3
	Bit 1

F1LX2
	Bit 0

F1LX1


F1Y2– F1X9 mask register

N & A(1) & F(17) with data bits as shown:

	Bit 7

F1LY2
	Bit  6

F1LY1
	Bit 5

F1LX14
	Bit  4

F1LX13
	Bit  3

F1LX12
	Bit 2

F1LX11
	Bit 1

F1LX10
	Bit 0

F1LX9


F1Y10– F1Y3 mask register

N & A(2) & F(17) with data bits as shown:

	Bit 7

F1LY10
	Bit  6

F1LY9 
	Bit 5

F1LY8
	Bit  4

F1LY7
	Bit  3

F1LY6
	Bit 2

F1LY5
	Bit 1

F1LY4
	Bit 0

F1LY3


F2X4 – F1Y11 mask register

N & A(3) & F(17) with data bits as shown:

	Bit 7

F2LX4
	Bit  6

F2LX3
	Bit 5

F2LX2
	Bit  4

F2LX1
	Bit  3

F1LY14
	Bit 2

F1LY13
	Bit 1

F1LY12
	Bit 0

F1LY11


F2X12 -  F2X5 mask register

N & A(4) & F(17) with data bits as shown:

	Bit 7

F2LX12
	Bit  6

F2LX11
	Bit 5

F2LX10
	Bit  4

F2LX9
	Bit  3

F2LX8
	Bit 2

F2LX7
	Bit 1

F2LX6
	Bit 0

F2LX5


F2Y6 – F2X13 mask register

N & A(5) & F(17) with data bits as shown:

	Bit 7

F2LY6
	Bit  6

F2LY5
	Bit 5

F2LY4
	Bit  4

F2LY3
	Bit  3

F2LY2
	Bit 2

F2LY1
	Bit 1

F2LX14
	Bit 0

F2LX13


F2LY14 – F2LY7 mask register

N & A(6) & F(17) with data bits as shown:

	Bit 7

F2LY14
	Bit  6

F2LY13
	Bit 5

F2LY12
	Bit  4

F2LY11
	Bit  3

F2LY10
	Bit 2

F2LY9
	Bit 1

F2LY8
	Bit 0

F2LY7


F3X8 – F3X1 mask register 

N & A(7) & F(17) with data bits as shown:

	Bit 7

F3LX8
	Bit  6

F3LX7
	Bit 5

F3LX6
	Bit  4

F3LX5
	Bit  3

F3LX4
	Bit 2

F3LX3
	Bit 1

F3LX2
	Bit 0

F3LX1


F3Y2– F3X9 mask register

N & A(8) & F(17) with data bits as shown:

	Bit 7

F3LY2
	Bit  6

F3LY1
	Bit 5

F3LX14
	Bit  4

F3LX13
	Bit  3

F3LX12
	Bit 2

F3LX11
	Bit 1

F3LX10
	Bit 0

F3LX9


F3Y10– F3Y3 mask register

N & A(9) & F(17) with data bits as shown:

	Bit 7

F3LY10
	Bit  6

F3LY9 
	Bit 5

F3LY8
	Bit  4

F3LY7
	Bit  3

F3LY6
	Bit 2

F3LY5
	Bit 1

F3LY4
	Bit 0

F3LY3


F4X4 – F3Y11 mask register

N & A(10) & F(17) with data bits as shown:
	Bit 7

F4LX4
	Bit  6

F4LX3
	Bit 5

F4LX2
	Bit  4

F4LX1
	Bit  3

F3LY14
	Bit 2

F3LY13
	Bit 1

F3LY12
	Bit 0

F3LY11


F4X12 -  F4X5 mask register

N & A(11) & F(17) with data bits as shown:
	Bit 7

F4LX12
	Bit  6

F4LX11
	Bit 5

F4LX10
	Bit  4

F4LX9
	Bit  3

F4LX8
	Bit 2

F4LX7
	Bit 1

F4LX6
	Bit 0

F4LX5


F4Y6 – F4X13 mask register

N & A(12) & F(17) with data bits as shown:
	Bit 7

F4LY6
	Bit  6

F4LY5
	Bit 5

F4LY4
	Bit  4

F4LY3
	Bit  3

F4LY2
	Bit 2

F4LY1
	Bit 1

F4LX14
	Bit 0

F4LX13


F4LY14 – F4LY7 mask register

N & A(13) & F(17) with data bits as shown:
	Bit 7

F4LY14
	Bit  6

F4LY13
	Bit 5

F4LY12
	Bit  4

F4LY11
	Bit  3

F4LY10
	Bit 2

F4LY9
	Bit 1

F4LY8
	Bit 0

F4LY7


The blocking of respective input works as follows:





If mask is set, then the respective input does not take part in the trigger logic, but allows the other inputs to decide the coincidence. Once mask is set the respective input is always treated as 1. On POWER and Z, all masks are set but no trigger will be generated. All inputs to be unmasked.Only ones which are no working should be masked.

Mask for intermediate trigger  

N & A(14) & F(17) with data bits as shown:

	Bit 7

F4Y
	Bit  6

F3Y
	Bit 5

F2Y
	Bit  4

F1Y
	Bit  3

F4X
	Bit 2

F3X
	Bit 1

F2X
	Bit 0

F1X


The blocking of respective intermediate trigger works as follows:


If mask is set, then the respective intermediate trigger does not take part in further trigger logic, but allows the other intermediate trigger inputs to decide the coincidence. Once mask is set the respective input is always treated as 0. On POWER and Z, all masks are set but no trigger will be generated.All intermediate triggers should be unmasked. Masking any one will not block the trigger logic.

Pulse width on front panel LTO and FTO:

N & A(15) & F(17) W(4-1)

Width in term of multiples of clock (50ns) specified by W in range 1 to 7. Default value is 6. The leading edge of the LTO and FTO signals starts with the leading edge of coincidence overlap (after the propagation delay involved in the trigger logic) and the width will vary from set to one additional clock jitter.

CAMAC_output

Multiplicity logic: All the inputs to be latched with LTO. Will be cleared with All multiplicity register clear .CLEAR (N & A(1) F(10))

           The input pattern can be read as follows:

Read Command N & A (0) & F (0) with data bits as shown

	Bit 15

F1LY2
	Bit 14

F1LY1
	Bit 13

F1LX14
	Bit 12

F1LX13
	Bit 11

F1LX12
	Bit 10

F1LX11
	Bit 9

F1LX10
	Bit 8

F1LX9


	Bit 7

F1LX8
	Bit 6

F1LX7
	Bit 5

F1LX6
	Bit 4

F1LX5
	Bit 3

F1LX4
	Bit 2

F1LX3
	Bit 1

F1LX2
	Bit 0

F1LX1


Read Command N & A (1) & F (0) with data bits as shown

	Bit 15

F2LX4
	Bit 14

F2LX3
	Bit 13

F2LX2
	Bit 12

F2LX1
	Bit 11

F1LY14
	Bit 10

F1LY13
	Bit 9

F1LY12
	Bit 8

F1LY11


	Bit 7

F1LY10
	Bit 6

F1LY9
	Bit 5

F1LY8
	Bit 4

F1LY7
	Bit 3

F1LY6
	Bit 2

F1LY5
	Bit 1

F1LY4
	Bit 0

F1LY3


Read Command N & A (2) & F (0) with data bits as shown

	Bit 15

F2LY6
	Bit 14

F2LY5
	Bit 13

F2LY4
	Bit 12

F2LY3
	Bit 11

F2LY2 
	Bit 10

F2LY1 
	Bit 9

F2LX14
	Bit 8

F2LX13


	Bit 7

F2LX12
	Bit 6

F2LX11
	Bit 5

F2LX10
	Bit 4

F2LX9 
	Bit 3

F2LX8 
	Bit 2

F2LX7 
	Bit 1

F2LX6 
	Bit 0

F2LX5 


Read Command N & A (3) & F (0) with data bits as shown

	Bit 15

F3LX8 
	Bit 14

F3LX7 
	Bit 13

F3LX6 
	Bit 12

F3LX5 
	Bit 11

F3LX4 
	Bit 10

F3LX3 
	Bit 9

F3LX2 
	Bit 8

F3LX1 


	Bit 7

F2LY14
	Bit 6

F2LY13
	Bit 5

F2LY12
	Bit 4

F2LY11
	Bit 3

F2LY10
	Bit 2

F2LY9
	Bit 1

F2LY8
	Bit 0

F2LY7


Read Command N & A (4) & F (0) with data bits as shown

	Bit 15

F3LY10
	Bit 14

F3LY9
	Bit 13

F3LY8
	Bit 12

F3LY7
	Bit 11

F3LY6
	Bit 10

F3LY5
	Bit 9

F3LY4
	Bit 8

F3LY3


	Bit 7

F3LY2
	Bit 6

F3LY1
	Bit 5

F3LX14
	Bit 4

F3LX13
	Bit 3

F2LX12
	Bit 2

F3LX11
	Bit 1

F3LX10
	Bit 0

F3LX9


Read Command N & A (5) & F (0) with data bits as shown

	Bit 15

F4LX12
	Bit 14

F4LX11
	Bit 13

F4LX10
	Bit 12

F4LX9
	Bit 11

F4LX8
	Bit 10

F4LX7
	Bit 9

F4LX6
	Bit 8

F4LX5


	Bit 7

F4LX4
	Bit 6

F4LX3
	Bit 5

F4LX2
	Bit 4

F4LX1
	Bit 3

F3LY14
	Bit 2

F3LY13
	Bit 1

F3LY12
	Bit 0

F3LY11


Read Command N & A (6) & F (0) with data bits as shown

	Bit 15

F4LY14
	Bit 14

F4LY13
	Bit 13

F4LY12
	Bit 12

F4LY11
	Bit 11

F4LY10
	Bit 10

F4LY9 
	Bit 9

F4LY8
	Bit 8

F4LY7 


	Bit 7

F4LY6 
	Bit 6

F4LY5 
	Bit 5

F4LY4
	Bit 4

F4LY3
	Bit 3

F4LY2 
	Bit 2

F4LY1 
	Bit 1

F4LX14
	Bit 0

F4LX13


Read Command N & A (7) & F (0) with data bits as shown

	Bit 15

0
	Bit 14 

0
	Bit 13

0
	Bit 12

0
	Bit 11

0
	Bit 10

 0
	Bit 9

0
	Bit 8

     0


	Bit 7

TO_F4Y
	Bit 6

TO_F3Y
	Bit 5

TO_F2Y
	Bit 4

TO_F1Y
	Bit 3

TO_F4X
	Bit 2

TO_F3X
	Bit 1

TO_F2X
	Bit 0

TO_F1X


Scalars: LTO, TO_F1X, TO_F2X, TO_F3X, TO_F4X, TO_F1Y, TO_F2Y, TO_F3Y, TO_F4Y. These are 16 bit counters (except LTO which is 24 bit counter) and will enable on CAMAC command N & A (1) & F (26) and will disable on CAMAC command N & A (1) & F (24)  readable on CAMAC bus using CAMAC commands as follows: 

Read Scalar_LTO_low -> N & A(14) & F(2)

Read Scalar_LTO_high -> N & A(15) & F(2)    (bits 8 - 15 are  0)

Read Scalar_Lam_width_low -> N & A(12) & F(2)

Read Scalar_LAM_width_high -> N & A(13) & F(2)    (bits 8 - 15 are  0)

Read Scalar_TO_F1X -> N & A(0) & F(2)

Read Scalar_TO_F2X -> N & A(1) & F(2)

Read Scalar_TO_F3X -> N & A(2) & F(2)

Read Scalar_TO_F4X -> N & A(3) & F(2)

Read Scalar_TO_F1Y -> N & A(4) & F(2)

Read Scalar_TO_F2Y -> N & A(5) & F(2)

Read Scalar_TO_F3Y -> N & A(6) & F(2)

Read Scalar_TO_F4Y -> N & A(7) & F(2)

These scalars can be cleared by individual commands or a common command as follows:

Reset Scalar_LTO -> N & A(14) & F(11)

Read Scalar_Lam_width-> N & A(12) & F(2)

Reset Scalar_TO_F1X -> N & A(0) & F(11)

Reset Scalar_TO_F2X -> N & A(1) & F(11)

Reset Scalar_TO_F3X -> N & A(2) & F(11)

Reset Scalar_TO_F4X -> N & A(3) & F(11)

Reset Scalar_TO_F1Y -> N & A(4) & F(11)

Reset Scalar_TO_F2Y -> N & A(5) & F(11)

Reset Scalar_TO_F3Y -> N & A(6) & F(11)

Reset Scalar_TO_F4Y -> N & A(7) & F(11)

Reset ScalarsAll   -> N & A(15) & F(11)

Total of 10 scalars. Scalers on LTO and LAM_width are 24bit. Others are 16bits.

	Bit 7

0
	Bit 6

0
	Bit 5

0
	Bit 4

0
	Bit 3

I
	Bit 2

LAM
	Bit 1

Q
	Bit 0

X


Status of control signal: The status of control signals like I, LAM, Q, X can be read on CAMAC command N & A(1) & F(1). The bit pattern is as follows,

General: Upon POWER_ON and command Z, the module will be in clear state i.e.

LAM_REQ will be cleared, LAM_ENABLE will be reset, scalers will be cleared 

Multiplicity Latched input registers will be zero. 

All input masks are set.

Command accepted signal X is generated on each implemented CAMAC command as listed in this manual.

Commands C and I have not been implemented and have no effect on the module. 

 CAMAC TRIGGER GENERATOR for INO, INO_CTG

Application_note

This note gives a guide on the use of the module. It describes the connector details, input interconnections and CAMAC commands.   It will serve as a ready reference.

Front panel diagram:

Connector details

There are two boards, one base board which has the CFMFC PC edge connector and one piggy board which connects to the base board via a 60 pin FRC cable. Connectors J1 and J2 on the piggy board, connectors J4 and J5 are on the base-board and connectors J3 on both the boards for interconnection.

	     J1
	      J2
	      J4
	     J5

	J1.1
	F1LX1+
	
	J2.1
	F2LX1+
	
	J4.1
	F3LX1+
	
	J5.1
	F4LX1+

	J1.2
	F1LX1-
	
	J2.2
	F2LX1-
	
	J4.2
	F3LX1-
	
	J5.2
	F4LX1-

	J1.3
	F1LX2+
	
	J2.3
	F2LX2+
	
	J4.3
	F3LX2+
	
	J5.3
	F4LX2+

	J1.4
	F1LX2-
	
	J2.4
	F2LX2-
	
	J4.4
	F3LX2-
	
	J5.4
	F4LX2-

	J1.5
	F1LX3+
	
	J2.5
	F2LX3+
	
	J4.5
	F3LX3+
	
	J5.5
	F4LX3+

	J1.6
	F1LX3-
	
	J2.6
	F2LX3-
	
	J4.6
	F3LX3-
	
	J5.6
	F4LX3-

	J1.7
	F1LX4+
	
	J2.7
	F2LX4+
	
	J4.7
	F3LX4+
	
	J5.7
	F4LX4+

	J1.8
	F1LX4-
	
	J2.8
	F2LX4-
	
	J4.8
	F3LX4-
	
	J5.8
	F4LX4-

	J1.9
	F1LX5+
	
	J2.9
	F2LX5+
	
	J4.9
	F3LX5+
	
	J5.9
	F4LX5+

	J1.10
	F1LX5-
	
	J2.10
	F2LX5-
	
	J4.10
	F3LX5-
	
	J5.10
	F4LX5-

	J1.11
	F1LX6+
	
	J2.11
	F2LX6+
	
	J4.11
	F3LX6+
	
	J5.11
	F4LX6+

	J1.12
	F1LX6-
	
	J2.12
	F2LX6-
	
	J4.12
	F3LX6-
	
	J5.12
	F4LX6-

	J1.13
	F1LX7+
	
	J2.13
	F2LX7+
	
	J4.13
	F3LX7+
	
	J5.13
	F4LX7+

	J1.14
	F1LX7-
	
	J2.14
	F2LX7-
	
	J4.14
	F3LX7-
	
	J5.14
	F4LX7-

	J1.15
	F1LX8+
	
	J2.15
	F2LX8+
	
	J4.15
	F3LX8+
	
	J5.15
	F4LX8+

	J1.16
	F1LX8-
	
	J2.16
	F2LX8-
	
	J4.16
	F3LX8-
	
	J5.16
	F4LX8-

	J1.17
	F1LX9+
	
	J2.17
	F2LX9+
	
	J4.17
	F3LX9+
	
	J5.17
	F4LX9+

	J1.18
	F1LX9-
	
	J2.18
	F2LX9-
	
	J4.18
	F3LX9-
	
	J5.18
	F4LX9-

	J1.19
	F1LX10+
	
	J2.19
	F2LX10+
	
	J4.19
	F3LX10+
	
	J5.19
	F4LX10+

	J1.20
	F1LX10-
	
	J2.20
	F2LX10-
	
	J4.20
	F3LX10-
	
	J5.20
	F4LX10-

	J1.21
	F1LX11+
	
	J2.21
	F2LX11+
	
	J4.21
	F3LX11+
	
	J5.21
	F4LX11+

	J1.22
	F1LX11-
	
	J2.22
	F2LX11-
	
	J4.22
	F3LX11-
	
	J5.22
	F4LX11-

	J1.23
	F1LX12+
	
	J2.23
	F2LX12+
	
	J4.23
	F3LX12+
	
	J5.23
	F4LX12+

	J1.24
	F1LX12-
	
	J2.24
	F2LX12-
	
	J4.24
	F3LX12-
	
	J5.24
	F4LX12-

	J1.25
	F1LX13+
	
	J2.25
	F2LX13+
	
	J4.25
	F3LX13+
	
	J5.25
	F4LX13+

	J1.26
	F1LX13-
	
	J2.26
	F2LX13-
	
	J4.26
	F3LX13-
	
	J5.26
	F4LX13-

	J1.27
	F1LX14+
	
	J2.27
	F2LX14+
	
	J4.27
	F3LX14+
	
	J5.27
	F4LX14+

	J1.28
	F1LX14-
	
	J2.28
	F2LX14-
	
	J4.28
	F3LX14-
	
	J5.28
	F4LX14-

	J1.29
	F1LY1+
	
	J2.29
	F2LY1+
	
	J4.29
	F3LY1+
	
	J5.29
	F4LY1+

	J1.30
	F1LY1-
	
	J2.30
	F2LY1-
	
	J4.30
	F3LY1-
	
	J5.30
	F4LY1-


	     J1
	      J2
	      J4
	     J5

	J1.31
	F1LY2+
	
	J2.31
	F2LY2+
	
	J4.31
	F3LY2+
	
	J5.31
	F4LY2+

	J1.32
	F1LY2-
	
	J2.32
	F2LY2-
	
	J4.32
	F3LY2-
	
	J5.32
	F4LY2-

	J1.33
	F1LY3+
	
	J2.33
	F2LY3+
	
	J4.33
	F3LY3+
	
	J5.33
	F4LY3+

	J1.34
	F1LY3-
	
	J2.34
	F2LY3-
	
	J4.34
	F3LY3-
	
	J5.34
	F4LY3-

	J1.35
	F1LY4+
	
	J2.35
	F2LY4+
	
	J4.35
	F3LY4+
	
	J5.35
	F4LY4+

	J1.36
	F1LY4-
	
	J2.36
	F2LY4-
	
	J4.36
	F3LY4-
	
	J5.36
	F4LY4-

	J1.37
	F1LY5+
	
	J2.37
	F2LY5+
	
	J4.37
	F3LY5+
	
	J5.37
	F4LY5+

	J1.38
	F1LY5-
	
	J2.38
	F2LY5-
	
	J4.38
	F3LY5-
	
	J5.38
	F4LY5-

	J1.39
	F1LY6+
	
	J2.39
	F2LY6+
	
	J4.39
	F3LY6+
	
	J5.39
	F4LY6+

	J1.40
	F1LY6-
	
	J2.40
	F2LY6-
	
	J4.40
	F3LY6-
	
	J5.40
	F4LY6-

	J1.41
	F1LY7+
	
	J2.41
	F2LY7+
	
	J4.41
	F3LY7+
	
	J5.41
	F4LY7+

	J1.42
	F1LY7-
	
	J2.42
	F2LY7-
	
	J4.42
	F3LY7-
	
	J5.42
	F4LY7-

	J1.43
	F1LY8+
	
	J2.43
	F2LY8+
	
	J4.43
	F3LY8+
	
	J5.43
	F4LY8+

	J1.44
	F1LY8-
	
	J2.44
	F2LY8-
	
	J4.44
	F3LY8-
	
	J5.44
	F4LY8-

	J1.45
	F1LY9+
	
	J2.45
	F2LY9+
	
	J4.45
	F3LY9+
	
	J5.45
	F4LY9+

	J1.46
	F1LY9-
	
	J2.46
	F2LY9-
	
	J4.46
	F3LY9-
	
	J5.46
	F4LY9-

	J1.47
	F1LY10+
	
	J2.47
	F2LY10+
	
	J4.47
	F3LY10+
	
	J5.47
	F4LY10+

	J1.48
	F1LY10-
	
	J2.48
	F2LY10-
	
	J4.48
	F3LY10-
	
	J5.48
	F4LY10-

	J1.49
	F1LY11+
	
	J2.49
	F2LY11+
	
	J4.49
	F3LY11+
	
	J5.49
	F4LY11+

	J1.50
	F1LY11-
	
	J2.50
	F2LY11-
	
	J4.50
	F3LY11-
	
	J5.50
	F4LY11-

	J1.51
	F1LY12+
	
	J2.51
	F2LY12+
	
	J4.51
	F3LY12+
	
	J5.51
	F4LY12+

	J1.52
	F1LY12-
	
	J2.52
	F2LY12-
	
	J4.52
	F3LY12-
	
	J5.52
	F4LY12-

	J1.53
	F1LY13+
	
	J2.53
	F2LY13+
	
	J4.53
	F3LY13+
	
	J5.53
	F4LY13+

	J1.54
	F1LY13-
	
	J2.54
	F2LY13-
	
	J4.54
	F3LY13-
	
	J5.54
	F4LY13-

	J1.55
	F1LY14+
	
	J2.55
	F2LY14+
	
	J4.55
	F3LY14+
	
	J5.55
	F4LY14+

	J1.56
	F1LY14-
	
	J2.56
	F2LY14-
	
	J4.56
	F3LY14-
	
	J5.56
	F4LY14-

	J1.57
	GND
	
	J2.57
	GND
	
	J4.57
	GND
	
	J5.57
	GND

	J1.58
	GND
	
	J2.58
	GND
	
	J4.58
	GND
	
	J5.58
	GND

	J1.59
	GND
	
	J2.59
	GND
	
	J4.59
	GND
	
	J5.59
	GND

	J1.60
	GND
	
	J2.60
	GND
	
	J4.60
	GND
	
	J5.60
	GND


CAMAC commands:

For INO_TEST module

Z,C reset module

F(26)A(0) Disable O/P

F(24)A(0) enable O/P

F(16)A(0) enable or Disable coincidence pattern out of 5 O/P

Example

COMMAND  F(24)A(0) 

Write DATA 1F

COMMAND F(16)A(0) will enable all 5 O/P 

                                                        
For INO module

Initialize, Z: Set LAM mask, clear LAM, clear multiplicity register, clear scalers

F(0)A(0-7): Read multiplicity registers

F(1)A(1): Read status

F(2)A(0-7): read scalers on intermediate coincidences

F(2)A(14-15): read LTO scaler

F(2)A(12-13): read LAM width scaler

F(8)A(0): Test LAM; Q=0 if LAM set

F(10)A(0): Clear LAM

F(11)A(0-15): clear scalers

F(17)A(0-13): set/reset input mask; 1 = set; 0 = reset

F(17)A(14): set/reset mask register for intermediate coincidence, 1 = set; 0 = reset

F(17)A(15): set LTO / FTO width

F(24)A(0): set LAM mask, disable LAM

F(26)A(0): reset LAM mask, enable LAM

Recommended sequence of commands:

After POWER_ON or any other time to reinitialize the module:

Initialize, Z

Set input mask registers

Enable LAM

After receiving LAM:

Read multiplicity registers

Clear multiplicity registers

Clear LAM

To estimate the system dead time



read LAM width scaler


Clear LAM width scaler

To estimate the input event rate 



Read trigger scalers



Clear trigger scalers

Sample programs for the CAMAC spy program

1. Enable all inputs, reset masks: for details of bit map, see user manual

Write data = 255

NAF = N 0 17 

Write data = 255

NAF = N 1 17 

Write data = 255

NAF = N 2 17 

Write data = 255

NAF = N 3 17 

write data = 255

NAF = N 4 17 

Write data = 255

NAF = N 5 17 

write data = 255

NAF = N 6 17 

Write data = 255

NAF = N 7 17 

Write data = 255

NAF = N 8 17 

Write data = 255

NAF = N 9 17 

Write data = 255

NAF = N 10 17 

Write data = 255

NAF = N 11 17 

Write data = 255

NAF = N 12 17 

Write data = 255

NAF = N 13 17 

Write data = 255

NAF = N 14 17 

Write data = 255

2. Disable all inputs, reset masks: for details of bit map see user manual

Write data = 0

NAF = N 0 17

Write data = 0

NAF = N 1 17 0

Write data = 0

NAF = N 2 17 0

Write data = 0

NAF = N 3 17 0

Write data = 0

NAF = N 4 17 0

Write data = 0

NAF = N 5 17 0

Write data = 0

NAF = N 6 17 0

Write data = 0

NAF = N 7 17 0

Write data = 0

NAF = N 8 17 0

Write data = 0

NAF = N 9 17 0

Write data = 0

NAF = N 10 17 0

Write data = 0

NAF = N 11 17 0

Write data = 0

NAF = N 12 17 0

Write data = 0

NAF = N 13 17 0

Write data = 0

NAF = N 14 17 0

3. Disabling a specific input: for ex. Input F3LX1

Write data = 254

 NAF = N 7 17 

This enables all the other inputs in this group

4. Enabling a specific input: for ex. Input F3LX1

Write data = 1

 NAF = N 7 17 

This disables all the other inputs in this group

If it is required to enable or disable a specific input without disturbing others in the group, then the previous value of the group must be preserved and only that bit should be modified. In the above example, if the group mask is defined as MASK8 then it should be modified as

MASK8 = MASK8 (bit wise AND) 0xFE to enable 

and

MASK8 = MASK8 (bit wise OR) 0x01 to disable

5. Reading multiplicity registers

NAF = N 0 26 ; enable LAM

Wait for LAM

NAF = N 0 0

Read data

NAF = N 1 0 

Read data

NAF = N 2 0

Read data

NAF = N 3 0

Read data

NAF = N 4 0

Read data

NAF = N 5 0

Read data

NAF = N 6 0

Read data

NAF = N 7 0

Read data

NAF = N 0 10; reset LAM

6. Reading scalers:

NAF = N 0 2

Read data

NAF = N 1 2 

Read data

NAF = N 2 2

Read data

NAF = N 3 2

Read data

NAF = N 4 2

Read data

NAF = N 5 2

Read data

NAF = N 6 2

Read data

NAF = N 7 2

Read data

NAF = N 12 2; read LAM width scaler low

Read data

NAF = N 13 2; read LAM width scaler high

Read data

NAF = N 14 2; read LTO scaler low

Read data

NAF = N 15 2; read LTO scaler high

Read data

CAMAC TRIGGER GENERATOR for INO, INO_CTG

Test setup and observations

Test setup:

As the inputs to the INO_CTG module are LVDS logic signal which are not available from any standard logic source, a special CAMAC module was developed to generate the required test signal. The highest fold of coincidence is 5 on 1fold signal from the first layer, thus this test module generates 5 separate signals. These five signals can be individually masked so as to facilitate the coincidence logic of the INO_CTG module. These five signals are in two groups of two and three signals. Two free running signal sources are built-in in the tset module with the width as well as overlap of these two signals groups settable using front panel pots. This allows the testing of the minimum overlap required to generate final trigger. The five signals can be individually connected to any of the 112 inputs using a specially made cable with 10 pin FRC connector on one end and five individual pairs on the other.

A block schematic of the test setup is as shown:



Step by step procedure for field testing of the module using CAMAC test module:

1. Adjust the widths and overlaps of the output signals of the test module as per required test stimulus.

2. Connect the five outputs to the required inputs of the INO_CTG module. 

3. Enable the test module by enabling and unmasking its five output 

4. Run test programs to test various features of the INO_CTG module as detailed in the application note.

5. Checking multiplicity logic:

a. To test two fold coincidence, first connect only one input to one of the four-fold input and check that no coincidence is generated. 

b. With this connection itself, mask-off the neighboring input so that this will now form a coincidence condition. Check if trigger is generated. LTO, FTO and LAM should be generated. Also the multiplicity registers read should be correct.

c. In similar way, check the two fold for other four-fold inputs.

d. Following similar steps check three, four and five fold coincidences.

6. In all of the above tests check if LAM enable, disable and reset works correctly. When LAM is set on trigger and is disabled, it will be removed from the CAMAC bus. But will be reasserted when it is enabled. Whereas resetting LAM will clear the LAM request register and enabling or disabling it will have no effect on the LAM on CAMAC bus. 

7. In all of the above tests, check if the respective scalers are read correctly.

8. Also check the LAM width scaler. It indicates the system dead-time. To test this feature, connect inputs so that trigger will happen. Enable the respective inputs. Wait for LAM, then put a varying delay before clearing LAM. Now read the LAM width scaler. With different values for the delay this should give appropriate value. 

Test program for checking LAM width scaler:


Wait for LAM


Delay (time value)



NAF = N 0 10 ; reset LAM


NAF = N 12 2


Read data


NAF = N 13 2


Read data

9. Set different values for LTO / FTO widths and check the output

Observations:

The module has been tested for all its functionality as listed as below and is found to be working as required.

Minimum overlap to detect coincidence

Propagation delay from input to output 
Feedback notes




LVDS





RECEI-VER











CAMAC bus





LAM logic





CAMAC registers





L   I 


V   n


D   p


S    u


      t


      s





Final Trigger Out





Trigger Out





Logic Trigger Out





Fig 1: Block Schematic of the INO_CTG





FPGA





Scaler_TO_F1X





5





5





One fold  Inputs, 14Y





One fold coincidence logic block











5





5





One fold  Inputs, 14X





input





MASK





L


O


G


I


C



























































































































































































































































































































































F(17)A(n)W(n)=1





Mask reg


Preset	     Q





Clr





To trigger logic 





Mask





Scaler_TO_F1X





One fold coincidence logic block











F(17)A(14)W(n)=1








F(17)A(14)W(n)=0


Or Z





F(17)A(n)W(n)=0


or Z





J1, J2, J3, J4, J5 are 60 pin standard FRC connectors with 0.1” pitch. 





COINCIDENCE BLOCK,





SCALERS, 





MULTIPLICITY








Mask reg


Preset	     Q





Clr





To trigger logic


Active high LTO





Mask





Intermediate trigger


     Active low





Piggy board





Base BoFrd





J1





J5





J2





J4





J3





J3

















 








_1211009999

_1211010041

_1211009958

